## Slide 1: A comparison study between Heart Rate obtained from Fingertip and Wrist position
Ayush Dhananjay Modani
MD Kamrul Hasan Foysal
Under Guidance of Dr. Jo Woon chong
Date : 29th November 2018

## Slide 2: Why heart rate monitoring is important??
Heart Diseases Problems rising day by day
807,775 deaths in just 2014
14% increase in deaths from (2009-2011) to (2012-2014)

## Slide 3: Heart Monitoring Devices

## Slide 4
ECG
Kardia
Pulse Oximeter
Smartphone
Smartwatch/Smartband

## Slide 5: Can the extraction position (place) in the body affects the heart rate??

## Slide 6: Lets compare the new technology (smartWatch and Smartphone) with the Golden standard

## Slide 7: Smartwatch/fitness band Heart Rate working

## Slide 8: Smartphones Herat rate working

## Slide 9: Methodology

## Slide 10: Flow Diagram
Start the Application
Capture the frames in Real-Time
Extract the Green Intensity from the images extracted
Save the intensity as a text file
Plot the intensity
Terminate Application after 2min
Apply Bit Arrangement 888-565
Apply Edge Detection and Smoothing
Use Curve Selection for PPG detection
Apply same for next Images
Extract Images from Video
Plot the PPG obtained
Place Components as shown and Start Recording
After 2min note the file name
Extract the file in MATLAB and process it
Filter the Signal using Moving Averaging Filter
Find the Peaks from the signal
Plot the Signal with Peaks
Calculate Heart Rate using Peaks and Average it
Intensity Method
Smartphone Application
Curve-based Method

## Slide 11: Result - PPG Signal
PPG from Nexus Device
PPG from Fingertip using Intensity Method
PPG from Wrist using Intensity Method
PPG from Fingertip using Curve-based Method

## Slide 12: Results – Heart Rate Comparison
Subject No | Age | Heart Rate from Fingertip |  | Heart Rate from Wrist |  | Heart Rate from NeXus
 |  | Heart Rate (IM) | Heart Rate (CM) | Heart Rate (IM) | Heart Rate (CM) | 
1 | 23 | 72 | 75 | 75 | 75 | 71.7757
2 | 27 | 64.2857 | 66.6666 | 62.0689 | 64.2857 | 66.2069
3 | 29 | 81.8182 | 85.7142 | 85.7143 | 81.8182 | 80.8421
4 | 21 | 58.0645 | 58.0645 | 60 | 62.0689 | 59.0769
5 | 19 | 62.0609 | 64.2857 | 60 | 60 | 60.9524
6 | 22 | 72 | 75 | 90 | 85.7143 | 73.1429
7 | 21 | 62.0689 | 62.0689 | 60 | 62.0689 | 61.44
8 | 18 | 81.8182 | 78.2608 | 78.2608 | 81.8182 | 80
9 | 26 | 72 | 75 | 72 | 72 | 73.8462
10 | 23 | 90 | 90 | 94.7368 | 90 | 90.3529
Chart: Heart Rate
Fingertip IM | 72 | 64.2857 | 81.8182 | 58.0645 | 62.0609 | 72 | 62.0689 | 81.8182 | 72 | 90
Fingertip CM | 75 | 66.6666 | 85.7142 | 58.0645 | 64.2857 | 75 | 62.0689 | 78.2608 | 75 | 90
Wrist IM | 75 | 62.0689 | 85.7143 | 60 | 60 | 90 | 60 | 78.2608 | 72 | 94.7368
Wrist CM | 75 | 64.2857 | 81.8182 | 62.0689 | 60 | 85.7143 | 62.0689 | 81.8182 | 72 | 90
NeXus | 71.7757 | 66.2069 | 80.8421 | 59.0769 | 60.9524 | 73.1429 | 61.44 | 80 | 73.8462 | 90.3529

## Slide 13: Results – Heart Rate Comparison
Chart: Heart Rate Comparison between Fingertip and NeXus
Intensity Method | 72 | 64.2857 | 81.8182 | 58.0645 | 62.0609 | 72 | 62.0689 | 81.8182 | 72 | 90
Curve-based Method | 75 | 66.6666 | 85.7142 | 58.0645 | 64.2857 | 75 | 62.0689 | 78.2608 | 75 | 90
NeXus | 71.7757 | 66.2069 | 80.8421 | 59.0769 | 60.9524 | 73.1429 | 61.44 | 80 | 73.8462 | 90.3529
Chart: Heart Rate Comparison between Wrist and NeXus
Intensity Method | 75 | 62.0689 | 85.7143 | 60 | 60 | 90 | 60 | 78.2608 | 72 | 94.7368
Curve-based Method | 75 | 64.2857 | 81.8182 | 62.0689 | 60 | 85.7143 | 62.0689 | 81.8182 | 72 | 90
NeXus | 71.7757 | 66.2069 | 80.8421 | 59.0769 | 60.9524 | 73.1429 | 61.44 | 80 | 73.8462 | 90.3529

## Slide 14: Results - Error
Subject No | NeXus Heart Rate | Error in Heart Rate from Fingertip |  |  |  | Error in Heart Rate from Wrist |  |  | 
 |  | Error in Heart Rate (IM) |  | Error in Heart Rate (CM) |  | Error in Heart Rate (IM) |  | Error in Heart Rate (CM) | 
 |  | In bpm | % | In bpm | % | In bpm | % | In bpm | %
1 | 71.7757 | -0.2243 | 0.3125 | -3.2243 | 4.49219 | -3.2243 | 4.49219 | -3.2243 | 4.49219
2 | 66.2069 | 1.9212 | 2.901812 | -0.4597 | 0.69434 | 4.138 | 6.250104 | 1.9212 | 2.901812
3 | 80.8421 | -0.9761 | 1.20742 | -4.8721 | 6.02669 | -4.8722 | 6.02681 | -0.9761 | 1.20742
4 | 59.0769 | 1.0124 | 1.713699 | 1.0124 | 1.713699 | -0.9231 | 1.56254 | -2.992 | 5.06459
5 | 60.9524 | -1.1085 | 1.81863 | -3.3333 | 5.46869 | 0.9524 | 1.562531 | 0.9524 | 1.562531
6 | 73.1429 | 1.1429 | 1.562558 | -1.8571 | 2.539 | -16.8571 | 23.0468 | -12.5714 | 17.1875
7 | 61.4400 | -0.6289 | 1.0236 | -0.6289 | 1.0236 | 1.44 | 2.34375 | -0.6289 | 1.0236
8 | 80.0000 | -1.8182 | 2.27275 | 1.7392 | 2.174 | 1.7392 | 2.174 | -1.8182 | 2.27275
9 | 73.8462 | 1.8462 | 2.500061 | -1.1538 | 1.56244 | 1.8462 | 2.500061 | 1.8462 | 2.500061
10 | 90.3529 | 0.3529 | 0.39058 | 0.3529 | 0.39058 | -4.3839 | 4.85197 | 0.3529 | 0.39058
Average Error |  |  | 1.570361 |  | 2.608522 |  | 5.481076 |  | 3.860297

## Slide 15: Results – Error Comparison
Chart: Error in Heart Rate between Wrist and Nexus
Intensity Method | 75 | 62.0689 | 85.7143 | 60 | 60 | 90 | 60 | 78.2608 | 72 | 94.7368
NeXus | 71.7757 | 66.2069 | 80.8421 | 59.0769 | 60.9524 | 73.1429 | 61.44 | 80 | 73.8462 | 90.3529
Chart: Error in Heart Rate between Fingertip and Nexus
Intensity Method | 72 | 64.2857 | 81.8182 | 58.0645 | 62.0609 | 72 | 62.0689 | 81.8182 | 72 | 90
NeXus | 71.7757 | 66.2069 | 80.8421 | 59.0769 | 60.9524 | 73.1429 | 61.44 | 80 | 73.8462 | 90.3529

## Slide 16: Result – Error comparison
Chart: Error Comparison - Intensity and Curve-based Method in Wrist
Intensity Method | 4.4922 | 6.2501 | 6.0268 | 1.5625 | 1.5625 | 23.0468 | 2.3437 | 2.174 | 2.5001 | 4.852
Curve-based Method | 4.4922 | 2.9018 | 1.2074 | 5.0646 | 1.5625 | 17.1875 | 1.0236 | 2.2728 | 2.5001 | 0.39058
Chart: Error Comparison - Intensity and Curve-based Method in Fingertip
Intensity Method | 0.3125 | 2.9018 | 1.2074 | 1.7137 | 1.8186 | 1.5626 | 1.0236 | 2.2728 | 2.5001 | 0.39058
Curve-based Method | 4.4922 | 0.69434 | 6.0267 | 1.7137 | 5.4687 | 2.539 | 1.0236 | 2.174 | 1.5624 | 0.39058

## Slide 17: conclusion
Heart Rate depends upon position from which data is obtained
Heart Rate from Fingertip is more Accurate than Heart Rate from Wrist
Intensity Method Heart Rate from Fingertip Images is having accuracy of 98%
Heart Rate from Curve-based Method using Fingertip Images is having accuracy of 97%
Intensity Method Heart Rate from Wrist Images is having accuracy of 94%
Heart Rate from Curve-based Method using Wrist Images is having accuracy of 96%

## Slide 18: References
1. Accuracy of smartphone apps for heart rate measurement, Thomas Coppetti, Andreas Brauchlin, Simon Muggler, European Journal of Preventive Cardiology 2017, Vol. 24(12) 1287–1293 The European Society of Cardiology 2017
2. https://www.mindmedia.com/de/produkte/nexus-10-mkii/
3. https://en.wikipedia.org/wiki/Ballistocardiography
4. The emerging field of mobile health Steven R. Steinhub, Evan D. Muse, and Eric J. Topo, HHS Public Access Author manuscript Sci Transl Med. Author manuscript; available in PMC 2016 February 10
5. Mobile Devices and Apps for Health Care Professionals: Uses and Benefits C. Lee Ventola, MS, P T. 2014 May; 39(5): 356–364., PMCID: PMC4029126, National Institutes of Health
6. Evidence assessing the diagnostic performance of medical smartphone apps: a systematic review and exploratory meta-analysis, Rahel Buechi,1 Livia Faes,2 Lucas M Bachmann,2 Michael A Thiel,1 Nicolas S Bodmer,2 Martin K Schmid,1 Oliver Job,1 Kenny R Lienhard3, BMJ Open 2017;7: e018280. doi:10.1136/ bmjopen-2017-018280
7. A. S. M. Mosa, I. Yoo, L. Sheets, "A systematic review of healthcare applications for smartphones", BMC Med. Inf. Decision Making, vol. 12, pp. 67, Jul. 2012.
8. Yadav, D.; Gupta, P.; Kaur, G.; Singh, J.; Gandh, I.M.; Singh, A. Implementing Edge Detection for Medical Diagnosis of A Bone in MATLAB. In Proceedinds of the International Conference on Computational Intelligence and Communication Networks (CICN), Mathura, India, 27–29 September 2013; pp. 270–274.
9. Zaman R, Cho CH, Hartmann-Vaccarezza K, et al. Novel fingertip image-based heart rate detection methods for a smartphone. Sensors. 2017;17(2):E358.
10. del Rosario MB, Redmond SJ, Lovell NH. Tracking the evolution of smartphone sensing for monitoring human movement. Sensors. 2015;15:18901–18933.
11. R. Zaman, C.H. Cho, K. Hartmann-Vaccarezza, T.N. Phan, G. Yoon, J.W. Chong, "Novel Fingertip Image-Based Heart Rate Detection Methods for a Smartphone", Sensors, vol. 17, no. 2, pp. 358, 2017
12. R. Zaman, C.H. Cho, Y. Kim, J.W. Chong, "A novel heart rate monitoring method using a smartphone", Proc. IEEE Healthcare Innovation Point-Of-Care Technologies (HI-POCT) Conference, pp. 5-8, 2016.
13. R.J. Duro, F.L. Pena, "Smartphone-based Photoplethysmogram Measurement" in Digital Image and Signal Processing Measurement Systems, Aalborg, Denmark:River Publisher, pp. 358, 2017.

## Slide 19: Demo

## Slide 20: Any questions?

## Slide 21: Thank You